

# Instructions for Use – Magic Strip™

**General instructions**

# •For the safety of the user and patient, it is highly recommended that only qualified Dentists work with the Magic Strip™

•Strauss’ Magic Strip™ is designed for several uses and are autoclavable.

•The strips are designed to work manually in the interproximal gap.

•Avoid touching soft tissue with the strips.

•Sterilize the Magic Strip™ before use.

**Sterilization instructions**

Instructions for cleaning and sterilizing the Magic Strip™

Disinfect and sterilize the strips before every use.

**Cleaning:**

•Clean the strips by dipping them in disinfection solution that contains a corrosion inhibitor for a minimum of 10 minutes.

•If needed -use brush to clean thoroughly.

•Rinse several times to remove all the disinfection solution.

**Cleaning with Ultrasound:**

**•**Strips may be ultrasonically cleaned.

•A cycle of 5 minutes is recommended, using a general-purpose cleaner.

•Rinse several times.

**Sterilizing the Magic Strip™ –Autoclave:**

•Place the strip in sterilization bags.

•Autoclave the strip for a minimum of 10 minutes at a temperature of 135ºC.

Following sterilizing or cleaning methods, dry the strips and store in a clean and moisture free environment.

**Note:**

•Use sterilizing and cleaning devices and materials according to the manufacturers recommended procedure.

•It is the responsibility of the user to ensure that sterilization is effective.